Architectural Specification

The audio hearing loop driver shall continuously deliver 2.4 A (RMS) when excited with a continuous 1 kHz sine wave. It shall cover an area of at least 377 sqft. (35 m²) with a room aspect ratio 3:1 according to IEC 60118-4:2006 standard.  The driver shall have a frequency response of at least 80 Hz to 6.5 kHz +/- 3 dB with THD+N less than 0.2%.  The unit shall incorporate audio processing to maximize the intelligibility for individuals with a t-coil equipped hearing aid. The unit shall incorporate metal loss correction to rectify frequency caused adjacent metal structures. The CLD1-XX Compact Loop Driver is specified.